## Slide 1: Portable Controlled Noise Environment
Jenny Liu, Amanda Spencer, Jennifer Wang
Team 34
Client: Dennis Barbour, M.D., Ph.D
Progress Presentation

## Slide 2: Need & Project Scope
Hearing loss affects 30 million Americans
Treat with hearing aids or cochlear implants
Design a portable controlled noise environment for an auditory training smartphone game to improve usability

## Slide 3: Specific Design Requirements
Physical Apparatus |  | 
Portable | Lightweight | <0.1 kg in ear (each ear), < 0.5 kg external
 | Compact | Fits in a 20x15x10x cm3 box
 | Intrinsic power source | Supports >90 minutes of use
Usable | Comfortable | 9/10 subjects can wear for > 90 minutes without discomfort.
 | Aesthetics | 9/10 subjects are willing to wear the device in public
Android Smartphone Interface | A/D sampling rate | > 40 kHz
 | Play output signal | > 128 kb/s
 | Control | Signal to noise ratio (SNR) as a double

## Slide 4: Specific Design Requirements
Sound Processing |  | 
Dynamic |  | <20 ms delay time
Ambient Noise Detection Range |  | 0-75 dBA SPL
Noise Attenuation Range |  | Up to 30 dBA
Frequency range |  | 20Hz – 20 kHz
Output | Speaker signal | Mixed signal with correct SNR, summing training signal and added noise
 | Range | 0-30 dB SPL
 | Frequency | 20 Hz – 20 kHz
 | Maximum | 130 dBA SPL

## Slide 5: Design Categories
Direct Plug-In
Telecoil stimulation
Portability
SNR Control
3D Speaker Array
Noise Cancellation
Headphone
Add-on
Introduce noise
Sound booth
LACETM
Train at home

## Slide 6: Design Choice – Headphone
 | Weight | Passive, No Speaker | Passive, With Speaker | Active and Passive, No Speaker | Active and Passive, With Speaker
Noise Cancelling Ability | 9 | 4 | 4 | 7 | 7
Ease of Assembly | 7 | 3 | 7 | 3 | 7
Total |  | 57 | 85 | 84 | 112
Notes: Fix…

## Slide 7: Design Modules
Modules
  Headphone
  External Microphone (9)
  Internal Microphone (2)
  Apparatus Attachment (8 )
  Apparatus body (5)
  Smartphone interface
  Control algorithm (3)
Design permutations = 2160

## Slide 8: External Microphones
Category
Dynamic
  Pros: rugged, long-lasting, less expensive
  Cons: frequency distortion
Condenser
  Pros: flatter frequency profile, better transient response
  Cons: power source, expensive
Directionality
Cardioid
Omni-directional
Placement
Boundary
Stalk

## Slide 9: External Microphones
Category
Dynamic
Condenser
Directionality
Cardioid
  Pros: resistant to feedback
  Cons: proximity effect
Omni-directional
  Pros: flatter frequency profile
  Cons: feedback
Placement
Boundary
Stalk
Flat

## Slide 10: Internal Microphone
Electret
  Pros: inexpensive, frequency profile flat
  Cons: larger, more inexpensive
MEMS
  Pros: very small, EM shielded
  Cons: frequency profile mostly flat

## Slide 11: Internal Microphone
Placement
  Behind the ear
  Behind the earlobe
  In external ear canal

## Slide 12: Internal Microphone
Placement
  Behind the ear
  Behind the earlobe
  In external ear canal

## Slide 13: Internal Microphone
Placement
  Behind the ear
  In external ear canal
  Behind the earlobe

## Slide 14: Apparatus Attachment
Clamp
Suction cup
Happeez adhesive
Dual Lock
Screw
Clasp
Snap-button
Glue
Notes: These solutions have a permanent aspect

## Slide 15: Apparatus Attachment
Clamp
Suction cup
Happeez adhesive
Dual Lock
Screw
Clasp
Snap-button
Glue

## Slide 16: Apparatus Attachment
Clamp
Suction cup
Happeez adhesive
Dual Lock
Screw
Clasp
Snap-button
Glue

## Slide 17: Internal Microphone Attachment
Bistable spring band
  Pros: off-the-shelf, storage shape
  Cons: difficult, hard to attach
  hard to control bending
Telescoping Metal Frame
Flexon
Bent rod
Glue
Magnet

## Slide 18: Internal Microphone Attachment
Bistable spring band
Telescoping Metal Frame
Flexon
  Pros: flexible, customizable, light, strong
  bio-compatible
  Cons: proprietary, unavailable
Bent rod
Glue
Magnet

## Slide 19: Software
Static Recording
  Pros: Absolute SNR control
  Cons: Does not account for NP

## Slide 20: Software
Static Recording
Iterative
  Adjust NT and S
  NP calculation
    Pros: Calculation has fewer steps (faster)
    Cons: Drift in NP may compromise SNR control
  SNR for negative feedback
    Pros: Precise SNR control
    Cons: More calculations increase delay

## Slide 21: Design Choice – External Microphone
 | Weight | Dynamic |  |  |  | Condenser |  |  |  | 
 |  | Cardioid |  | Omnidirectional |  | Cardioid |  | Omnidirectional |  | 
 |  | Stalk | Flat | Stalk | Flat | Stalk | Flat | Stalk | Flat | Boundary
Frequency fidelity | 10 | 3 | 2 | 5 | 4 | 8 | 7 | 10 | 9 | 10
Rapid Transient Response | 8 | 8 | 8 | 8 | 8 | 10 | 10 | 10 | 10 | 10
Low Noise | 10 | 10 | 10 | 9 | 9 | 8 | 8 | 7 | 7 | 7
Avoid feedback | 10 | 10 | 10 | 10 | 10 | 6 | 6 | 6 | 6 | 6
Rugged | 7 | 10 | 10 | 10 | 10 | 8 | 8 | 8 | 8 | 7
Low Weight | 5 | 8 | 8 | 8 | 8 | 7 | 7 | 7 | 7 | 5
Size | 6 | 8 | 8 | 8 | 8 | 7 | 7 | 7 | 7 | 3
Aesthetically Acceptable | 5 | 2 | 7 | 2 | 7 | 2 | 7 | 2 | 7 | 6
Total |  | 244 | 259 | 254 | 269 | 285 | 300 | 295 | 310 | 305

## Slide 22: Design Choice – Internal Microphone
 | Weight | Electret Condenser | MEMS | Dynamic
EM Shielded | 7 | 7 | 10 | 5
Frequency Fidelity | 10 | 10 | 9 | 5
Small Size | 10 | 6 | 10 | 0
Low Noise | 10 | 7 | 7 | 10
Rugedgness | 5 | 6 | 5 | 10
Sum |  | 309 | 355 | 235

## Slide 23: Design Choice – Apparatus Attachment
 | Weight | Bistable | Telescoping | Flexon | Bent Rod | Glue | Magnet
Customizable fit several headphone styles | 3 | 6 | 6 | 8 | 8 | 10 | 10
Comfortable | 7 | 8 | 8 | 8 | 8 | 10 | 10
Non-interfering | 10 | 8 | 8 | 8 | 8 | 0 | 0
Stable/Secure stay > 90 minutes | 10 | 4 | 6 | 10 | 10 | 8 | 8
Available on market | 10 | 10 | 4 | 0 | 6 | 10 | 10
Total |  | 294 | 254 | 260 | 320 | 280 | 280

## Slide 24: Design Choice – Internal Microphone Attachment
 | Weight | Clamp | Suction Cup | Happeez Adhesive | Dual Lock | Screw | Clasp | Snap Button | Glue
Reversability | 3 | 10 | 10 | 10 | 8 | 8 | 7 | 8 | 3
Ruggedness | 9 | 8 | 3 | 5 | 8 | 9 | 8 | 7 | 9
Non interfering | 8 | 5 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Comfort | 8 | 5 | 8 | 8 | 8 | 8 | 8 | 8 | 8
Aesthetics | 6 | 4 | 7 | 7 | 7 | 7 | 7 | 7 | 7
Non- destructive | 10 | 3 | 8 | 10 | 10 | 5 | 7 | 6 | 7
Total |  | 236 | 323 | 361 | 382 | 341 | 349 | 333 | 346

## Slide 25: Chosen Design
Headphone: commercial
External microphone: omnidirectional electret (condenser)
Internal microphone: MEMS, positioned behind earlobe
Attachment method: Dual Lock
Bent wire attachment: stainless steel, beryllium, titanium
Smartphone Interface:
  Two 3.5 mm port adapter
  USB

## Slide 26: Chosen Design
Headphone: commercial
External microphone: omnidirectional electret (condenser)
Internal microphone: MEMS, positioned behind earlobe
Attachment method: Dual Lock
Bent wire attachment: stainless steel, beryllium, titanium
Smartphone Interface:
  Two 3.5 mm port adapter
  USB

## Slide 27: Winner
Internal Microphone
External Microphone

## Slide 28: Roles
Tasks | Jenny | Amanda | Jennifer
Contact Manager | X |  | 
Weekly reports | X | X | X
Website Maintenance |  | X | X
Hardcopy Manager |  | X | 
Background - Literature | X | X | 
Background - Patents | X |  | X
Mechanical |  | X | X
Software Interfacing | X |  | X
Software Development | X |  | X
Sound Acquisition | X | X | 
Signal Processing | X |  | 
Sound Production |  | X | X

## Slide 29: Calendar
 | = due date
 | = working

## Slide 30: Calendar
Tasks/ Milestones | 5-Nov | 7-Nov | 12-Nov | 14-Nov | 19-Nov | 21-Nov | 28-Nov | 30-Nov | 3-Dec | 5-Dec | 11-Dec
Team Formation |  |  |  |  |  |  |  |  |  |  | 
Project selection |  |  |  |  |  |  |  |  |  |  | 
Weekly Reports |  |  |  |  |  |  |  |  |  |  | 
Background Research |  |  |  |  |  |  |  |  |  |  | 
Patents and Prior Art Research |  |  |  |  |  |  |  |  |  |  | 
Develop Project Specifications |  |  |  |  |  |  |  |  |  |  | 
Prelim. Oral Report |  |  |  |  |  |  |  |  |  |  | 
Prelim. Written Report |  |  |  |  |  |  |  |  |  |  | 
Web Page Development |  |  |  |  |  |  |  |  |  |  | 
Design Options Developement |  |  |  |  |  |  |  |  |  |  | 
Design Options Analysis |  |  |  |  |  |  |  |  |  |  | 
Prog. Oral Report |  |  |  |  |  |  |  |  |  |  | 
Prog Written Report |  |  |  |  |  |  |  |  |  |  | 
Design Refinement |  |  |  |  |  |  |  |  |  |  | 
Software Development |  |  |  |  |  |  |  |  |  |  | 
Peer Review |  |  |  |  |  |  |  |  |  |  | 
Final Oral Report + Demo |  |  |  |  |  |  |  |  |  |  | 
Final Written Report |  |  |  |  |  |  |  |  |  |  | 
Poster Design |  |  |  |  |  |  |  |  |  |  | 
 | = due date
 | = working

## Slide 31: Questions?
WW2 UK Acoustic Mirror (Pre-RADAR)